Korea Basic Science Institute COMMODITY DESCRIPTION
Commodity	French Pressure Cell
HSK No.	8479.82.9000	Page page of pages	Date	2014.02.12
Description					Unit	Q'ty
A. Features 1. Air powered, French press style, ultra high pressure homogenizer 2. Pneumatically pressure controlled 3. Batch type & quiet operation 4. Simple, small size, bench top & multi-pass capability 5. No electrical connection required 6. Fully clean-in place (CIP) sterilizable 7. Suitable for clean room and GMP facility 8. Wetted parts are autoclavable 9. No o-ring or gasket use in entire path of the product (No leakage) 10. Ready to use and easy to install B. System Configuration French Pressure Cell press C. Specification - 10ml or more sample volume - 3500 bar 40,000 psi or more - Low power requirement - Maximum operating pressure Volumes / flow rate : System can handle 5-20ml sample at pressures more of 40,000psi in single shot and larger volumes uo to many litres in multi-cycle mode. Made of operation selected by operator. - Minimum practical sample size - Safety The high pressure system is totally enclosed and interlocked to prevent operation unless the cover door is closed. - Homogenising Valves : The homogenising valve features ceramic needle and seat with profiled geometry for optimised performance. - Pressure & Temperature Measurement : A strain gauge pressure measurement system with digital pressure display is provided. Pressure measurement is through scaling of hydraulic drive pressure. - System is simply operated through push button control on cell/intensifier cycle and manual adjustment of disruption/homogenisation pressure. - Sterilizable cell & valve D. Standard Acc's 1) Ccell body 2) Outlet cooler 3) Cell body heat exchanger 4) After cooler / heat exchanger 5) Refrigerated bath circulator 6) Maintenance kit E. Remarks 1. Two year warranty after installation should be guaranteed by bidder. 2. Installation, test-run and operation training should be performed by a qualified engineer. 3. Operation manual and service manual should be provided.					sys set ea ea ea ea ea ea	1 1 1 1 1 1 1 1